## Sayfa1
YÖNTEMLER
TOTAL ANTIOXDANT STATUS (TAS) (mmol/L)
TAS levels were measured using commercially available kits (Relassay, Turkey). The novel
automated method is based on the bleaching of characteristic color of a more stable ABTS
(2,2 ′ - Azino-bis(3-ethylbenzothiazoline-6-sulfonic acid)) radical cation by antioxidants. The
assay has excellent precision values, which are lower than 3%. The results were expressed as
mmol Trolox equivalent/L (Erel O. A novel automated direct measurement method for total
antioxidant capacity using a new generation, more stable ABTS radicalcation. Clin Biochem
2004;37:277-85.)
(Marka: Relassay)
TOTAL OXIDANT STATUS (TOS) (µmol/L)
TOS levels were measured using commercially available kits (Relassay, Turkey. In the new
method, oxidants present in the sample oxidized the ferrous ion-o-dianisidine complex to
ferric ion. The oxidation reaction was enhanced by glycerol molecules abundantly present in
the reaction medium. The ferric ion produced a colored complex with xylenol orange in an
acidic medium. The color intensity, which could be measured spectrophotometrically, was
related to the total amount of oxidant molecules present in the sample. The assay was
calibrated with hydrogen peroxide and the results were expressed in terms of
micromolar hydrogen peroxide equivalent per liter (μmol H2O2 equivalent/L). ( Erel O. A
new automated colorimetric method for measuringtotal oxidant status. Clin Biochem
2005;38:1103-11. ).
(Marka: Relassay)
OXIDATIVE STRESS INDEX (OSI)
The ratio of TOS to TAS was accepted as the oxidative stress index (OSI). For calculation, the
resulting unit of TAS was converted to μmol/L, and the OSI value was calculated according to
the following Formula : OSI (arbitrary unit) =
TOS (μmol H2O2 equivalent/L) / TAC (μmol Trolox equivalent/L). (1-3).
1. Yumru M, Savas HA, Kalenderoglu A, Bulut M, Celik H, Erel O. Oxidative imbalance in
bipolar disorder subtypes: a comparative study. Prog Neuropsychopharmacol Biol Psychiatry.
2009 Aug 31;33(6):1070-4.
2. Kosecik M, Erel O, Sevinc E, Selek S. Increased oxidative stress in children exposed to
passive smoking. Int J Cardiol 2005;100:61–4.
3. (Harma M, Harma M, Erel O (2003) Increased oxidative stress in patients with
hydatidiform mole. Swiss Med Wkly 133:563-536).
Paraoxonase-1 PON-1) U/L
Measurement of paraoxonase activity;
Paraoxonase activity was measured using
commercially available kits (Relassay, Turkey).
The rate of paraoxon hydrolysis (diethylpnitrophenylphosphate)
was measured by monitoring the increase of
absorption at 412 nm at 37 °C. The amount of generated p-nitrophenol
was calculated from the molar absorption coefficient at pH 8.5, which
was 18.290 M−1 cm−1 Paraoxonase activity was expressed as U/L serum
(Marka: Relassay)
Myeloperoxidase (MPO) U/L
MPO posseses various catalytical activities. It exhibits the main catalytical activity by the
production of hypochlorous acid (HClO) from hydrogen peroxide (H2O2) and chloride anion,
Cl- (or halide). MPO also exhibits peroxidase activity that catalyzes oxidation of a number of
substrates by H2O2. These reactions categories have been widely used to assess the
activities of MPO.
The Relassay Myeloperoxidase Chlorination Activity Assay Kit and The Relassay
Myeloperoxidase Peroxidation Activity Assay Kit are quantitative and colorimetric assay kits
for measuring the myeloperoxidase activity within a sample. In the The Relassay
Myeloperoxidase Chlorination Activity Assay Kit, MPO catalyzes the formation of
hypochlorous acid, which reacts with taurine to form taurine chloroamine. Taurine chloroamine reacts with the chromophore TNB, resulting in the formation of the colorless
product DTNB. One unit of MPO activity is defined as the amount of enzyme that hydrolyzes
the substrate and generates taurine chloramine to consume 1.0 μmole of TNB per minute.In
the The Relassay Myeloperoxidase Peroxidation Activity Assay Kit, MPO catalyzes odianisidine
to colored o-dianisidyl radical using H2O2. The increasing absorbance is
monitored at 412 nm and the activity is measured kinetically. This kit can be used manually
and easily adapted to automated analyzers.
(Marka: Relassay)
Super Oxide Dismutase (SOD) U/ml
The role of speroxide dismutase is to accelerate the dismutation of the toxic radical, produced
during oxidative energy processes to hydrogen peroxide and molecular oxygen. This method
employs xanthine and xanthine oxidase to generate superoxide radicals which react with 2-(4-
iodophenyl)-3-(4-nitrophenol)-5-phenyltetrazolium chloride to form a red formazan dye.. the
superoxide dismutase activity is then measured by the degree of inhibiton of this reaction
(Marka: Relassay)
Calcium mg/dl
Calcium in the sample reacts with arsenazo III forming a coloured complex that
can be measured by spectrophotometry.
(Otto Scientific)
HDL Cholesterol mg/dl
Enzymatic colorimetric test
• Sample and addition of R1
• Addition of R2 and start of reaction
In the first step LDL, VLDL and Chylomicrons are eliminated and transformed to
non reactive compounds and specific condition for the reaction. By the second
reagent only the HDL-Cholesterol is subject to color reaction
Cholesterol Esterase
Cholesterol ester + H2O Cholesterol + fatty acid
Cholesterol Oxidase
Cholesterol + O2 Cholesten-3-on + H2O2
Peroxidase
H2O2 + phenol + 4-aminoantipyrine quinoneimine dye+4 H2O
(Otto Scientific)
LDL Cholesterol mg/dl
İlk adımda, HDL, VLDL ve Şilomikronlar elimine edilir ve reaksiyon için özel koşulda reaktif olmayan bileşiklere dönüştürülür. İkinci reaktif sadece LDL-kolesterol renk reaksiyonudur.
Kolesterol esteraz
Kolesterol ester + H2O kolesterol + yağ asidi
Kolesterol Oksidaz
Kolesterol + O2 kolesten-3-on + H2O2
peroksidaz
H2O2 + fenol + 4- aminoantipirin kinon boyası +4 H2O
(Otto Scientific)
Triglycerides mg/dl
Triglycerides in the sample originates, by means of the coupled reactions described below, acoloured complex that can be measured by spectrophotometry.
Triglycerides + H2O lipase Glycerol + Fatty acids
Glycerol + ATP glycerol kinase Glycerol – 3 – P + ADP
Glycerol – 3 –P + O2 G-3-P-oxidase Dihidroxyacetone – P +H2O2
2 H2O2 + 4 – Aminoantipyrine + 4 – Chlorophenol G-3-P-oxidas Quinoneimine + 4 H2O
(Otto Scientific)
Phosphorus mg/dl
Inorganic phosphate forms an ammonium phosphomolybdate complex having the formula(NH4)3[PO4(MoO3)12] with ammonium molybdate in the presence of sulfuric acid.
The complex is determined photometrically in the ultraviolet region (340 nm).
(Otto Scientific)
Chicken Insulin
INS ELISA kit applies the competitive enzyme immunoassay technique utilizing a polyclonal anti-INS antibody and an INS-HRP conjugate.
The assay sample and buffer are incubated together with INS-HRP conjugate in pre-coated plate for one hour. After the incubation period,
the wells are decanted and washed five times. The wells are then incubated with a substrate for HRP enzyme. The product of the enzyme-substrate reaction
forms a blue colored complex. Finally, a stop solution is added to stop the reaction, which will then turn the solution yellow.
The intensity of color is measured spectrophotometrically at 450nm in a microplate reader.
The intensity of the color is inversely proportional to the INS concentration since INS from samples and INS-HRP
conjugate compete for the anti-INS antibody binding site. Since the number of sites is limited, as more sites are occupied by
INS from the sample, fewer sites are left to bind INS-HRP conjugate. A standard curve is plotted relating the intensity of the color (O.D.)
to the concentration of standards. The INS concentration in each sample is interpolated from this standard curve.
(Marka: Relassay)
Micro Protein
Bradford yöntem
(Marka: Relassay)