## Sayfa1
YÖNTEMLER
Zinc (ZN) μg/dl
Zinc found in the samples change the red-orange color of 5-Br-PAPS to light pink under alkaline conditions.
The change of absorbance at 548 nm is proportional to total zinc level in the sample.
The assay can be calibrated with zinc sulfate dissolved in deionized water.
(Relassay, Turkey)
Copper (CU) μg/dl
Copper found in the samples change the red-orange color of DiBr-PAESA to violet under acidic conditions.
The change of absorbance at 572 nm is proportional to total copper concentration in the sample.
The assay can be calibrated with copper sulfate dissolved in deionized water.
(Relassay, Turkey)
Iron (FE) (μg/dl)
Colorimetric assay pH < 2,0 Transferrin-Fe-complex apotransferrin + Fe3+ ascorbate Fe3+ Fe2+ FerroZine® + Fe2+ colored complex Under acidic conditions,
iron is liberated from transferrin. Lipemic samples are clarified by the detergent. Ascorbate reduces the released Fe3+ ions to Fe'2+ ions which react with
FerroZine to form a colored complex. The color intensity is directly proportional to the iron concentration and can be measured photometrically.
Mangan
Yöntem: ICP-MS
Selenyum analizinde kullandığımız cihaz ; Agilent 7700 ICP-MS
Kullanılan iç kalite kontroller; Recipe clinchek control for trace elements
Saf standart; Selenium Standard for ICP-MS Sigma-Aldrich